Bed Entrapment Zone Measuring Tool

Prior to measuring the zones, take the following person measurements. If an adult’s head or neck measurements are smaller than these Standard measurements or an adult’s chest depth is larger than 318mm, consider implications for bed safety.

| Head Breadth(distance across the face from ear to ear): Greater than 120mm | Neck breadth (diameter/width) Greater than 60mm | Chest Depth (anterior to posterior of chest) Less than 318mm |
| --- | --- | --- |

For instructions on measuring the zones with regard to bed rails, refer to Instructions for Measuring Entrapment Zones

Some zones may be relevant for other bed equipment /accessories where gaps are created in the bed system and are indicated as follows: BS (Bed Stick), M (mattress)

| Entrapment / Falls Zone | Test Tool(s) Required | Equipment Program Requirements for Adults | Equipment Program Requirements for Children | Pass (P) or Fail (F) |
| --- | --- | --- | --- | --- |
| Zone 1 - Within the rail BS | Cone, spring scale | <120mm | <30mm |  |
| Zone 2 - Under the rail | Cone, spring scale | <120mm with compressed mattress | <30mm with compressed mattress |  |
| Zone 3 - Between the rail and mattress BS | Cone | <120mm | <30mm |  |
| Zone 4 - Under the rail at the end of rail | Cone and cylinder (assembled) | <60mm and opening under the rail at its end >60⁰ | Full length rails supplied |  |
| Zone 5 – Between split rails | N/A | N/A – these rails not supplied | N/A – these rails not supplied |  |
| Zone 6 - Between the end of the rail and the side edge of the head or footboard | Cylinder | Head end: <60mm and Foot end: <60mm or >318mm | <30mm |  |
| Zone 7 - Between the head or footboard and the mattress M | Cone | <120mm | <30mm |  |
| Height to top side of rail from uncompressed mattress | Tape measure | 220mm | Case-by-case |  |

# Text boxes

This is a modified version of SA Health’s comparison chart

Zone 7

Zone 6

Zone 4

Zone 5

Zone 3

Zone 1

Zone 2